# CELL SORTING PROCEDURE

What to do:

- Schedule a meeting to discuss your experiment before planning your sort

- Be on time- If you bring more samples than originally planned for or if you are late, your sort will only go until the next appointment arrives.

- If you need to cancel your sort, please give 24 hours notice. This allows someone else to schedule for that time slot. If you do not cancel within 2 hours of your sort, you will be charged a cancellation fee.

- A 15-minute set-up fee will be included for all sorts.

- Bring proper controls. A negative or unstained control is required per cell type. Single-stained controls are necessary for multi-color sorts. Each control should be suspended at an approximate concentration of 1 million cells per 500ul (minimum).

- All samples, including controls, need to be filtered before sorting. Please ask for filters if you need them.

- Please bring collection media and extra buffer.

What to know:

- Due to the length of time required to sort cells, multiple sorts or a large number of cells may require an early morning or late evening appointment.

- Sample concentration (2-10 x 10^6/mL)

Smaller cells should be suspended at a concentration of approximately 5-10 million cells per milliliter. Larger cells need to be somewhat more dilute, 3-5 million cells per milliliter usually works well.

- Adherent cells- require special care for sorting

One major problem with adherent cells in flow cytometry are aggregates. It is essential that the samples be in a single-cell suspension. Not only can aggregates cause clogging of the sorter, they may also pose a much worse problem for the investigator. For example, if there is an aggregate of two or more cells containing a positively stained cell, this event may be seen by the cytometer as a positive single event. If the event is not screened out by the scatter gates or by another parameter such as propidium iodide fluorescence for live cell/dead cell discrimination, the aggregate will be sorted. Thus, any negative cells that are stuck to the single positive cell will be sorted as passenger cells and will be co-cultured with the “sorted” populations.

- Appropriate sample buffers for cell sorting

Most suspension cells do poorly at high pH. Since carbonate buffers are rather weak and allow the sample pH to increase in air (where sorting takes place), we recommend that the sample buffer be either PBS additionally buffered with 25mM HEPES (at a pH appropriate for the cells), or a culture media without bicarbonate buffers (also buffered with 25mM HEPES) or phenol red. Indicators such as phenol red (used in tissue culture media) and other compounds, such as certain B vitamins, will adversely affect the cytometer’s measurement of specific cell fluorescence, and should be avoided.

- Avoid indicators such as PHENOL RED

Carefully choose the appropriate sample buffer for cell sorting. See recipe below.

# Basic Sorting Buffer

- 1x Phosphate Buffered Saline (Ca/Mg++ free)

- 1mM EDTA

- 25mM HEPES pH 7.0

- 1% Fetal Bovine Serum (Heat-Inactivated)

***Conditions that can adversely impact your cell sample prior to sorting include cell media, sheath fluid, buffering capacity and protein concentration. Low cellular viability, auto fluorescence, and cell aggregates will result in poor sorts. Identification of these factors and taking the appropriate action to remedy the problem will improve sort purity, yield, and cell viability.